# Polarcam recording guide

Draft for review • 16 October 2026

For frame stacks and individual spot recordings on the prepared lab PC. Before your first session, ask Denny to show you the illumination, focus and stage-drive controls.

Start the GUI

- Close IDS peak Cockpit if it is open. In the polarcam folder, double-click Launch-Polarcam.cmd. Keep the console window open while using the GUI.

- Set Experiment to hinges or bfm. Check the recording folder shown at the top. The default is data/<experiment>/polarcam/YYYY-MM-DD. Use Recording folder… only if you need a different location; select the folder for the whole session.

- Set Recording to the sample type. Add a Filename label, e.g. DNA_W1_F01_wide. In notes, record the sample/well, wide or narrow field, illumination settings and anything unusual. Update these before each capture.

- In Live video, click Start live feed. Start with Exp (ms) = 1.2 and Gain = 20, then Apply. Exposure and gain are shared across tabs. Click a rod to centre the Magnifier and focus the sample. Display grayscale stretch changes the display only; it does not make the saved signal brighter.

Capture a matching background

- Have the interference drive set and checked before capturing background. Keep that setting during measurements and put its frequency and amplitude in the notes. Set Drive (manual) to its actual On/Off state; this dropdown records a note and does not control the hardware.

- Keep the laser, other illumination, exposure and gain as intended for the measurement. Click Capture background. Move the sample continuously sideways in X/Y throughout the 6-second capture, so rods move across the image. Keep Z fixed.

- Wait for capture to finish and the BG name and ID to update. Tick Subtract background. This also ticks Save with background subtraction; check that it remains ticked in the recording tab.

- Capture a new background after changing illumination path/power, exposure, gain, focus or drive settings. In particular, use a matching background when switching between wide and narrow field. Check the BG name before recording.

Record frame stacks

- In Spot analysis, choose Detection → Intensity for hinges or other rods that may be stationary; use Time variation for rotating BFM rods. Start with 150 Frames, leave FPS at its initial setting, and click Capture frame stack. Wait for saving and analysis to finish.

- Check Spot brightness — no display gain. The 0–255 scale uses the brightest pixel across the captured frames. 250 and above is near the limit. Aim for clear spots without clipping the ones you want to measure; reduce illumination if necessary, then repeat the background and trial capture.

- Click several marked spots and use Play spot to inspect them. Check that they are isolated, stay in the window and look distinct from the surrounding background. Detection marks candidates; it can miss dim rods or select background features. Do not use clipped spots for angle measurements.

- For each new field, move the sample in X/Y, update the field number and notes, then repeat the capture. Keep comparable sample/control settings. Note any repeated field or deliberate settings change. Each stack is saved automatically in frames.

Record an individual spot

- In Live video, restart the feed if needed and click an isolated rod. Click Use magnifier centre for Spot examine. Give it a spot label, e.g. DNA_W1_S01_wide, and record why you selected it.

- In Spot examine, check exposure, gain, BG name and Save with background subtraction. Start with ROI size = 14 sensor pixels; enlarge the window if the rod does not fit with some margin. Use the same ROI size and FPS request for comparable recordings; the camera may deliver a lower FPS.

- Click Start recording, record a short trial of about 10 seconds, then Stop recording. Wait for analysis to finish. Check the rod stays in the preview window and the signal is not clipped before making a longer recording. Recordings save automatically in spots; Save copy… is only for an extra copy.

Finish

- Wait for all recording and saving to finish, then Stop live feed and close the GUI. Back up the whole dated recording folder, including backgrounds, NPY data, JSON notes and raw-peak files. Follow the bench shutdown procedure shown by Denny.

For analysis, detector changes, APD recordings or modifications to the GUI, ask Denny or use the README and code.